MSTARI 3/96 for pre-processed data files

from the LKB-Gel Scanner

General: MSTARI is a program for the interactive evaluation of sedimentation equilibrium data derived from the Rayleigh interference optics of an Analytical Ultracentrifuge. In this implementation, the original mainframe FORTRAN version [] has been completely rewritten in MS-QUICKBASIC 4.5 for an IBM compatible PC. It allows the model independent evaluation of sedimentation equilibrium data with the M* function [], the conventional way via the local slopes of the ln J vs. r2 plots to give point average molecular weights Mw,app. (r) as well as in this implementation the analysis of the data set with the OMEGA function [], a function particularly well suited for the characterization of associative or interacting systems. The radial fringe shift must be determined in advance either by on- or off-line data capture techniques. MSTARI works with ASCII data files derived from off-line data capture with an LKB 2202 laser densitometer (Bromma, Sweden). The data from the gel scanner have been processed with a Fourier transform algorithm "ANALYSER" to give the relative radial fringe shift j(r) []. If other data capture systems are used, it should be no principal problem to transform the data files into the appropriate format for MSTARI so long as a dataset j(r) vs. r is provided. However, if the XL-I Rayleigh interference optics will be available, it is likely that a version of MSTARI will be launched to directly read in these data files. Two example files named Chitosan & B31i can be found as example files on the disk which have been used for this documentation.

The structure of the ASCII data file which is read into MSTAR is as follows:

166,  168	   |First No. = number of data points		Second No. = Data point number of cell bottom

5.96891, 0.00000	|First column				Second column

5.96997, 0.00418	|Radial position			relative fringe shift

5.97103, 0.00434

5.97209, 0.02925

5.97315, 0.02891

5.97421, 0.02896

.

.

.

6.13237, 4.18981

6.13343, 4.33798

6.13449, 4.51114

6.13555, 4.73505

6.13661, 4.95714

6.13768, 5.10715

6.13874, 5.37861

6.13980, 5.69965

6.14086, 5.92260

6.14192, 6.17202

6.14298, 6.49248

6.14404, 6.79060

6.14617

First the number of data points, indicated in the first row is read in. Then the radial position of the cell bottom is calculated with its data point number and the radial increment between the data points.

Program description:

To run the program, you need an IBM compatible PC with VGA graphics. Some time consuming calculations are carried out during the evaluation. Therefore the applictaion of at least an 80386 equipped with a math coprocessor is recommended. Nevertheless, MSTARI will still run on an 8086 PC. For the data printout, a HP Deskjet printer is recommended. In contrast to earlier program versions, the MSTAR program has been extended so much, that the OMEGA part of the program had to be written as a separate program. Hence there are two programs called: NMSTARI.EXE and OMEGAI.EXE. Each of this programs can be executed by typing its name without the extension .EXE when in the appropriate directory. However, to ensure a more convenient data analysis especially of several data files, these evaluation programs have been embedded into a menue program called START.EXE. It is recommended to use this program to switch between the different evaluation procedures. To use the START.EXE program, a directory C:\MSTAR must be created and the files: START.EXE, NMSTARA.EXE, NMSTARI.EXE, OMEGAA.EXE and OMEGAI.EXE be copied from the disk to this directory. Now the program is started by changing into C:\MSTAR and typing START. The following menue will appear:

Now, choose menue point 3. For the OMEGA analysis, choose point 4. The OMEGA analysis is subject of a separate manual. After MSTARI is started, you will see the introductory screen which asks you, if you want to switch the printer on or off. The current MSTARI version was written for the HP-Desk Jet printers. It might run with laser printers as well but it does not with any matrix printer. If you have a HP Deskjet printer connected to your computer, you will get a printout of all graphs you select while you are evaluating. Hence you are able to individually create your data documentation. You must make shure that your Deskjet printer is on-line because the program will terminate if any printer error occurs

If the program asks for a user input, it generally only accepts the given possibilities regardless if the user input is in lower or upper case. If any other input is given, the question for the user input is repeated. After the decision about the printouts, the user gets into the data input screen. Here, all relevant parameters are requested. The requested run number and the comment are only for documentation purposes in the printouts and appear on the title of the data printout.

After the data input, corrections can be made on the following screen by typing the appropriate number. If the displayed data are correct, proceed with <ENTER>.

Now the program requests the name of the data file. It is important to give the full name of the data file including the correct path and extension because the program will terminate if it cannot find the data file. Therefore you need to confirm that your data file input is correct before you can proceed.

If a physically meaningless negative j(r) value is detected while reading in the data, this data point is set to 0. On the next screen, your data file name is requested. You need to give the full pathname and the filename (less than 8 digits) without extension for your data files. It is recommended to use the same filename (without extension) as has been used for the ASCII data file which is currently evaluated. Therefore this filename is displayed. Again, you need to reconfirm your filename. Afterwards you are asked if you want a display of a graph of the relative fringe shift in dependence on the radial position. This is the data as read in from the ASCII file.

The data output of the MSTARI program has no publication quality and serves only documentation purposes. As it also might be desired to carry out further evaluations with the data generated during the evaluation, it was decided to store all relevant data in ASCII files. This enables the user to read the data with every program, self written or commercial. The ASCII files generated by MSTARI are classified by their file extension as follows with a = Meniscus:

.IDJ	j(r) vs r, Relative fringe shift j(r)  vs radial position

.IJA	J/(r2 - a2) vs jrr2−a2	Data of plot for Ja determination

.J	J(r) vs r, Absolute fringe shift J(r)  vs radial position

.IJ	ln J vs x, ln J vs x, x = (r2 - a2) / (Bottom2 - a2)

.IM	M* vs x, M* function vs x

.IMW	Mw,app.vs x, Apparent point weight average molar mass vs x

.IWJ	Mw,app.  vs J, Apparent point weight average molar mass vs J

.IMZ	Mz,app. vs x, Apparent point z-average molar mass vs x

.IZJ	Mz,app. vs J, Apparent point z-average molar mass vs J

.MYR	Plot of the Roark-Yphantis molar mass My2,r vs r

.MYJ	Plot of the Roark-Yphantis molar mass My2,r vs J

.MPR	Plot of the predicted molar mass Mpred. vs J

These data files only consist of the x-y ASCII data of the corresponding plots. It is highly recommended to use the printer output of MSTARI to create a printed documentation. This documentation can then serve as reference for the ASCII files created by MSTARA. If for example a data file B31M.TXT has been evaluated and the name for the MSTARI ASCII data files was chosen to be B31M, the data file for the ln J vs r2 plot is B31M.IJ This open structure of MSTARI enables for example the comparison of files from different evaluations with a plot program. One application of this can be the overlay plot of Mw,app. vs J plots for different solute concentrations to distinguish between polydispersity and self-association. Other applications the data files created by MSTARI might be useful for are the fitting of the data to different models by means of regression techniques.

On the next screen the relative fringe shift j(r) or also delta J in dependence of the radius is displayed if this has been selected by the user.

On the next screen, the user is asked to perform a Ja extrapolation (Ja is the fringe shift at the meniscus). This extrapolation works with a graphics cursor and uses a plot of J/(r2 - a2) vs jrr2−a2 as above. If the user wants this, a screen is displayed, where the operation of the graphics cursor is explained.

This extrapolation via the determination of the data range with a graphics cursor is used several times in the MSTARI program and is explained here in general. The user can decide which range of data he wants to use for the extrapolation by moving the yellow graphics cursor with the cursor keys. The actual x-position of the graphics cursor in pixel is displayed in a window in the upper right hand corner. If the left end of the data range is marked with the yellow graphics cursor, the position needs to be fixed and is then displayed as green line. After it is confirmed by the user that the selected position is O.K., the right end of the data range needs to be selected in analogy. The data between the marks 1 & 2 are now selected for the extrapolation. It is possible to select several more data ranges. After the whole data range has been selected for the extrapolation, the question for another mark needs to be negotiated. Now the user can input the degree n (up to 15) of a regression polynomial for the extrapolation. The regression polynomial is of the type: y=i=0nai⋅xi. Generally, the higher degree polynomials require more calculation time. After the calculation, the regression polynomial is displayed in red together with the white data points. It may happen that the polynomial does not fit the data very well or that the extrapolated value is not likely. In such a case, the whole extrapolation procedure inclusive the selection of the data range needs to be repeated. For the Ja extrapolation, the user needs to define a linear range or ranges of data points near the meniscus. Therefore the user is not asked for the input of the degree of the regression polynomial here as the default is a linear regression.

It is generally possible to rescale each axis of each plot. If this is desired, press any key after the plot has been displayed and input the axis to be rescaled as well as the new minimum and maximum values. The previous values are displayed to give an orientation. The rescaling function can be used as as often as the user wants to. Due to certain limitations of the programming language Quick Basic 4.5, the numbers of the graph cannot be located at the position they should appear on the axis in every case. But it should still be obvious which number belongs to which tick mark.

As soon as the user has selected the linear data range(s) near the meniscus, the question for another mark needs to be negotiated. Now the calculation is started and the regression line is displayed in red together with the data points. The extrapolated Ja value is also displayed which is: 2 intercept / slope of the regression line [2]. This extraploation can be repeated as often as desired.

As numerous procedures are known to work out Ja [], it is recommended to reconfirm the previously determined Ja value. MSTARI offers an alternative method which employs a sliding strip procedure to work out the local slopes in the j/(r2 - a2) vs jrr2−a2 plot. This procedure is very time consuming and can be skipped if the user wishes to. However, modern 486 or Pentium PC's already allow a rapid calculation of the sliding strip procedure. 4 sliding strip lengths representing different degrees of data smoothing are employed. The 4 graphs are displayed in different colours and allow an estimation of Ja. In the chosen example, the Ja of about 0.02 can be confirmed. If the user prefers another method to determine Ja, he/she can still use this because on the next screen, he/she can either input a Ja value or keep the Ja value of the first extrapolation.

Once Ja has been evaluated is this the way the routine can now use the absolute fringe concentration J(r) = j(r) + Ja. From the absorption gradient and the radial positions of the meniscus and cell bottom, the cell loading concentration J0 (in absolute fringe units) can be predicted according to:

J0=2Bottom2−Meniscus2MeniscusBottomJ(r)dr

On the next screen, the second diagnostic plot of the MSTARI program, a plot of ln J vs x, the normalized squared radial position, is displayed.

If the ln J vs x plot bends downwards in the region of the cell bottom (x = 1), the system is nonideal, if it bends upwards it is either polydisperse or self-associating. Warning: Both effects can cancel each other out and lead to a straight plot. In some cases this can be misleading.

The next plot displays the MSTAR function as a function of x. M* is calculated via [1]:

M∗J(r)−JakJar2−a2+2karrJ(r)−Jadr

with k = 1−vρw22RT, v= partial specific volume of the solute, r = buffer density, w = angular velocity, T = thermodynamic temperature, R = gas constant. First, only the MSTAR plot is displayed to give the possibility of a printout without any regression polynomial. As usual the plot can be scaled to taste. In the next MSTAR plot, the yellow graphics cursor must be used to define the range(s) of data for the extrapolation of M* to the cell bottom (x = 1) as it is described above. After the selection of a regression polynomial and its calculation, the M* plot is displayed with the regression polynomial in red. The value of the M* function at the cell bottom (x = 1) gives the apparent weight average molar mass Mw,app. of the solute [2]. This value is displayed. In most cases it is very useful to perform several extrapolations to get an impression of the error range of Mw,app. Two notes of caution: (1) M* data in the first half of the cell (x < 0.5) are already noisy because of the low fringe increments. (2) Sometimes it occurs that M* bends very suddenly strongly in the region of the cell bottom. In such a case, the strong sudden bend (a cell base effect) should not be taken into account. If there are no data points in the region of the cell bottom, the extrapolation becomes more difficult. In these cases it is recommended to use regression lines (degree 1) or other low degree polynomials. The following screens show a variety of different MSTAR extrapolations on the same dataset.

When Mw,app.. has been estimated to your satisfaction, you can the proceed to obtain derived point weight average molecular weights from the local slopes in the ln J vs. r2 resp. x plot via:

Mw,app.(r)=dlnJ(r)kd(r2)

As the determination of Mw,app. (and later of Mz,app.) might be time consuming depending on the computer used, the user if he/she wishes can skip this evaluation. In that case the whole evaluation is finished and the user will be asked if he wants to perform another evaluation.

Experience has shown, that it is necessary to use a so called sliding strip procedure to smooth the noise in the data out a bit which has a significant influence on the quality of the determined Mw,app.(r). The sliding strip procedure works the local slope of a data point out by determining the slope of a regression polynomial of either degree 1 or 2 (This can be selected) at this point. And a further warning: Errors in Ja effect point average Mw(r)app. data more than Mw,app. data. The length of the sliding strip determines the number of data points on each side of the point of interest, which are taken into account for the polynome regression. The greater the number of data points, the greater the smoothing and the calculation time. MSTARI recommends a quadratic polynomial and a sliding strip length which is the number of the data points devided by 10. However, the sliding strip length should in any event be between 10 and 20. When pressing <Enter> the default resp. recommended values are used.

The next two screens display respectively graphs of the apparent point weight average molar mass Mw,app.(r) versus the radial position x and the fringe concentration J which is proportional to the solute concentration. It can be seen that the Chitosan which has been chosen for this example obviously is nonideal as it can be expected for a charged polysaccharide. This could already be detected with the downward bend in the ln J vs x plot. The plots of the apparent point weight average molar mass are useful diagnostic plots to detect polydispersity, nonideality or self-association as well as the detection of different solute species e.g. the smallest and biggest species. Furthermore this classical evaluation gives a further way of evaluating Mw,app. which can be very useful if the M* extrapolation to the cell bottom is very difficult or even impossible. Furthermore it is possible to obtain the smallest sedimending species by extrapolating the Mw,app.(r) on 0 concentration (J = 0) in many cases. Therefore it is recommended to perform a calculation of Mw,app.(r) in every case.

After the two Mw,app. screens have been displayed, the user is asked if he/she wants to perform the evaluation of the apparent point z-average molar mass Mz,app.(r). As this evaluation is based on a double differentiation of the ln J vs x data, it is not very reliable in many cases.

Mz,app.(r)=Mw,app.(r)+1kdd(r2)ln1rdJ(r)dr

For the determination of Mz,app.(r) the same sliding strip settings as for the Mw,app.(r) determination are used. In analogy to Mw,app.(r), two plots of Mz,app.(r) are obtained: one vs x and one vs J.

In the example shown, despite the noise - an unavoidable feature unless very heavy smoothing is employed - one can still make out a clear trend in the data from higher to lower molecular weights towards the cell bottom due to the nonideality.

From Mz,app.(r) and Mw,app.(r), one is able to calculate the Roark-Yphantis molar mass My,2(r) [] which is independent from first order non-ideality effects and the predicted molar mass Mpred.(r):

My,2(r)=Mw,app.(r)2Mz,app.(r)

MPred.(r)=Mw,app.(r)[3Mw,app.(r)−2Mz,app.(r)

As discussed above, Mw,app.(r) and especially Mz,app.(r) are extremely sensitive to experimental noise. Therefore the quality of the My,2(r) plots depends on the quality of the Mw,app.(r) and Mz,app.(r) plots.

The same is valid for the plots of Mpred.(r). As can be seen in the presented example, the inaccuracies caused by the differentiation process for Mw,app.(r) and especially Mz,app.(r) can be found for Mpred. as a dependent value.

After these plots have been displayed, the user can get a printout of the numerical data, if the printer has been switched on in the beginning of the program. On the last screen, the user is asked if he wants to perform another evaluation. If no is selected, the program is terminated, if yes is selected, all data are cleared from the memory and another evaluation can be done. This can for example be an OMEGA analysis on the same data set.

Acknowledgements:

I would like to thank Stephen E. Harding in which group this work was carried out for the encouragement to undertake this project and for helpul discussions.

Contact adress:

If you should detect any bugs or other errors in this program, please contact me so that I can change the program in a further version. Please contact me also if you have suggestions how to improve this program, especially concerning the information which can be derived from the data set. If possible, please use e-mail.

Helmut Cölfen

Max-Planck-Institute for Colloid and Interface Research

Colloid Chemistry Department

Kantstr. 55

D-14513 Teltow

Germany

Fax: 49-3328-46215

e-mail: Coelfen@mpikg-teltow.mpg.de

It is also possible to contact:

Stephen E. Harding

National Centre for Macromolecular Hydrodynamics

University of Nottingham

Sutton Bonington LE12 5RD

England

Fax: 44-602-516142

e-mail: sczsteve@szn1.nott.ac.uk

References

[?]	S. E. Harding, J. C. Horton, P. J. Morgan in S. E. Harding, A. J. Rowe & J. C. Horton (Editors) "Analytical Ultracentrifugation in Biochemistry and Polymer Science", The Royal Society of Chemistry, Cambridge, UK (1992) 275

[?]	J. M. Creeth, S. E. Harding; J. Biochem. Biophys. Methods 7 (1982) 25

[?]	B. K. Milthorpe, P. D. Jeffrey, L. W. Nichol; Biophys. Chem. 3 (1975) 169

[?]	A. J. Rowe, S. Wynne Jones, D. G. Thomas, S. E. Harding in S. E. Harding et. al (Editors) "Analytical Ultracentrifugation in Biochemistry and Polymer Science", The Royal Society of Chemistry (1992) 49

[?]	J. M. Creeth, R. H. Pain; Progr. Biophys. Mol. Biol. 17 (1967) 217

[?]	D. Roark and D. Yphantis, Ann. N. Y. Acad. Sci. 164 (1969) 245