TWOCOMP - A suite of programs for simulating and modeling

sedimentation equilibrium in solutions containing two

macromolecular components

Allen P. Minton (minton@helix.nih.gov)

National Institutes of Health

Version 1.0 - February 18, 1998

The programs included in this package may be used to model

sedimentation equilibrium data acquired on solution mixtures of

two interacting macromolecular components (A and B).  The theory

and assumptions underlying the analysis are described fully in

A.P. Minton, "Alternative strategies for the characterization of

associations in multicomponent solutions via measurement of

sedimentation equilibrium", Prog. Colloid Polym. Sci. 107: 11-19;

1997. It is strongly recommended that this paper be read and

understood in order to obtain optimal results from the programs

included here, and, importantly, to understand the limitations of

the analysis and of these programs.  Reprints may be obtained

upon request from the author.

The file TWOCOMP.EXE is a self-decompressing archive which should

be placed in a newly created subdirectory and executed, at which

time, the following decompressed files should appear:

Read.me - a copy of this file

2compsim.exe - a simulation program

plateau.exe - a program for determining loading absorbances

2c1sfit.exe - a program for modeling sed. eq. absorbance

gradients acquired at a single wavelength

2c2sfit.exe - a program for modeling sed. eq. absorbance

gradients acquired at two wavelengths

The following data are necessary.

1. Sedimentation equilibrium data on solutions of each of the

two components individually.  These are to be analyzed by

conventional one-component sedimentation equilibrium modeling to

ascertain the buoyant molar mass of each component, and also to

detect the presence of self-association or heterogeneity in each

component.  If heterogeneity is encountered in either component,

the analysis of hetero-association cannot proceed.  The presence

of reversible self-association to dimers (A2 and/or B2) can be

accommodated within the model for hetero-association.  These

experiments and analyses should be completed before proceeding to

study mixtures of A and B.  The absorbance spectrum of each

component should be recorded, and two wavelengths selected

(lambda1 and lambda2) so that the ratio of absorbance of A at

lambda1 and lambda2 is significantly different from that of B.

It is also preferable that the wavelengths be selected so that

they are located at or near absorbance maxima or minima, or at

least in a region of the spectrum where the absorbance is not

changing sharply with wavelength.

2. Sedimentation equilibrium data should be acquired on three

solution mixtures of A and B having different compositions.  The

loading concentrations of each component in each solution (in

g/l) must be recorded and furnished as input to the modeling

program.  After loading each solution into the cells,

rotor, and centrifuge, radial scans at lambda1 and lambda2 should

be taken at low speed (1 or 2 Krpm) immediately after starting

the centrifuge.  The program PLATEAU may then be used to

determine the loading absorbance of each sample at each

wavelength.  These loading absorbances must be recorded and

furnished as input to the modeling program.  PLATEAU may also be

used to determine the radial position of the meniscus of each

sample, which must be supplied to the modeling program.

3. Following centrifugation to sedimentation equilibrium, each

sample is scanned at the two wavelengths lambda1 and lambda2.  If

you want to be extra thorough, you may want to change rotor speed

and centrifuge to sedimentation equilibrium at the second speed.

Consistency of results obtained at the two speeds is an important

test of the assumption that true sedimentation equilibrium has

been achieved.

4. The mixture data may now be analyzed using the programs

2C1SFIT or 2C2SFIT.  The former program utilizes data acquired at

one wavelength only, whereas the latter program utilizes data

obtained at both wavelengths, and is in principle capable of

achieving higher resolution --- but only if the data are good

enough and all of the underlying assumptions are valid.

5. The model parameters include the buoyant molar masses of

each of the components, and association constants for formation

of homodimers of each of the individual components (MA, MB,

logK20, and logK02).  In principle these can be determined from

the mixture data.  But if you have done the one-component studies

properly, you can safely constrain these parameters to their

predetermined values, thereby greatly improving the resolution of

the subsequent analysis of the association scheme and the

determination of best-fit values of association constants.  The

most efficient strategy for analysis of heteroassociation that we

have employed is as follows:

First constrain all heteroassociation constants to be zero

(i.e., logKij = -20).  If the two proteins are significantly

heteroassociating under the conditions of your experiments, then

the global fit should be poor.  Then gradually "turn on" one

heteroassociation constant at a time, i.e., allow logK11 to

float, and use an initial guess of ca 0.2 (small but finite).  If

that doesn't provide a good fit, turn it off (constrain it to a

large negative value) and turn on another constant, say logK21.

If a good fit is not obtained with any single constant, try

turning on more than one constant.  Perhaps you have multiple

association states present.  Also, try different starting guess

values.  Because these models are nonlinear, sometimes the

function minimizer can get trapped in a local minimum that has an

unacceptably high value of chi-square and prevents it from

finding a global minimum that has a much lower and quite

acceptable value of chi-square.  If none of the combinations

work, then there may be higher association states that we haven't

yet put into the model, or the absorbance doesn't obey Beers'

law, or there was some irreversible aggregation and the system

didn't obey conservation of mass.

If at first you don't succeed ...

Modeling is an art, not an exact science.  The computer can help

you, but it can't do the job alone.  Insight, and most

importantly, well-designed experiments and RELIABLE DATA are

essential.  The program 2COMPSIM included here will allow you to

simulate sedimentation equilibrum data obtained from associating

systems, including pseudo-random noise comparable to that

characteristic of a reasonably well-maintained XLA.  Using these

simulated data you can test your modeling strategy and explore

the information content of your proposed experiments.  To

simulate data obtained at multiple wavelengths, use the same

concentrations, rotor speeds, etc., but change the values of the

extinction coefficients (epsA and epsB) to reflect the different

specific absorbances of each component at the different

wavelengths.

Comments, questions and constructive suggestions may be addressed

to minton@helix.nih.gov.

Good luck!  -- Allen Minton

###############################################################################################################################################################################################################################################################################################################################################################################################